Early detection of patient deterioration at home
using smart medical sensors
Bachelor Thesis
Julian Lobbes

---
Agenda
Background, Motivation, Objectives, Tasks
Medwings Software
Usage Trial
Discussion of Results
Demonstration

---
Background: Deterioration
Adverse clinical events are preceeded by a decline in vital functions
Clinical deterioration
Observable in advance

---
Background: Early warning scores
Early Warning Scores (EWS)
indicator for:
decline in overall vital functions
severity of illness
predictor for:
length of hospital stay
ICU admission
Mortality

---
Background: Early warning scores
Widely used internationally
In hospitals and ambulant care
For triage, monitoring, care escalation and communication

---
Background: Modified Early Warning Score (MEWS)

---
Motivation: Remote Patient Monitoring (RPM)
Recent improvements in smart monitoring technology have opened new avenues

---
State of the Problem
Can EWS and RPM be combined?
Can we use IoT for that?
Can patients do this themselves?

---
Research Questions
Can a Remote Warning Score Monitoring (RWSM) system, feasible for everyday use, be implemented using smart medical devices commercially available today?
What are the technical and operational challenges of implementing such a system?
Can existing, validated EWSs be utilized in RWSM?

---
Research Tasks
Select an EWS
Select suitable smart devices
Design and develop an RWSM software system
Trial the system in real world conditions
Evaluate

---
EWS Selection
MEWS
Heart Rate
Blood Pressure (systolic)
Body Temperature
Respiratory Rate
AVPU Score (Glasgow Coma Scale)

---
Device Selection
Criteria:
Commercially available
Clinical approval
Non-intrusive, mobile use
Programmatic access to data

---
EWS Selection
MEWS
Heart Rate
Blood Pressure (systolic)
Body Temperature
Respiratory Rate
Shortness of breath
Blood oxygen saturation (SPO2)
AVPU Score (Glasgow Coma Scale)

---
Medwings: Requirements
Medwings
(Mobile Early Deterioration Warning System)
Software Requirements:
Security
Availability
Usability
Validity

---
Medwings: Architecture

---
Medwings: Architecture

---
System Trial: Methodology
Daily usage for one week
Five notifications per day at three hour intervals
Consistent measurement procedure
Data gathered:
Vitals data, MEWS data
Measurement environment (at home/on the go)
Subject reactivity
Connection quality
Measurement failures

---
System Trial: Measurement Failure Classification

---
System Trial: Analysis
How well did the subject adhere to the schedule?
Measurements missed (P1 failures)
Time between notification and interaction?
How effective was the system?
How many measurements, how many MEWS calculations?
Were MEWS alerts triggered correctly?
How frequent were measurement failures?

---
System Trial: Measurement and Failure Counts
7 days × 5 notifications = N = 35 interactions recorded
26 at home, 9 on the go
Patient missed 7 measurements
5 of those on the go
28 measurements made
17 successful MEWS claculations
11 synchronization timeouts
Almost all caused by BPM Core or Thermo

---
System trial: Measurement and Failure Counts

---
System trial: Number of Aborted Measurements Per Device

---
System trial: Connection Quality & Synchronization Failures

---
System Trial: User Reactivity
Average delay between notification and user taking measurement:
36 minutes at home
68 minutes on the go
Average time taken to carry out all measurements:
4.5 minutes (both at home and on the go)

---
Discussion
Software requirements were met
Medwings was successful in retrieving a wide range of vital measurements regularly
System was used successfully at home and on the go
Abnormal vital readings were detected and reflected by MEWS

---
Discussion
Only 61% of notifications resulted in MEWS records
Many synchronization failures
No correlation with connection quality
Caused by sync delays on the Withings Devices/Cloud
Watch was fine, BPM Core and Thermo problematic
Many skipped measurements on the go
Limited access to vitals devices
Taking measurements manually not always feasible

---
Limitations
MEWS algorithm not used in its original form
Only a single, biased test subject
Short trial period
Fairly complex setup process for end users

---
Future Work & Improvements
Dependence on third party API introduces many problems
Synchronization delays
Privacy and security issues
Risk of system becoming unusable if manufacturer ceases operations
Skipping cloud API would be beneficial
Continuous monitoring would be better

---
Conclusion
RWSM is feasible using modern IoT equipment
Not quite with existing EWSs
May hold considerable value in:
Reducing health system resources
Devices are cheaper, reduces medical staff workload
Improving clinical outcomes for patients at risk
Earlier detection of adverse events
Improving patient’s quality of life
Improved autonomy and mobility

---
Conclusion
More research should be done
Better iterations of software like Medwings are realistic
Especially promising considering the development of IoT health devices

---
Discussion